Urinary Tract Infections Chromogenic Medium

| Product No. | Product Category | Specification |
| --- | --- | --- |
| CRM022B | Dehydrated Culture Medium | 500g/bottle |

Intended Use

For isolation and identification of urinary tract pathogens.

Principle and Interpretation

Peptoneand yeast extract supply essential nutrients and vitamins in this medium. Chromogenic mixdetermines the color of different pathogens in genus level. Agar is coagulant.

Formulation

| Ingredients | /liter |
| --- | --- |
| Peptone | 12.0 g |
| Yeast extract | 5.0 g |
| Chromogenic mix | 1.0 g |
| Agar | 15.0 g |
| pH 7.0±0.2 at 25°C |  |

Preparation

Suspend 33.0g in 1 L distilled or deionized water. Heat with frequent agitation and boil to completely dissolve the powder. Distribute into flasks. Autoclave at 121℃ for 15 minutes.

Quality Control

Cultural characteristics observed after incubation at 35-37°C for 18-24 hours

| Quality control strains | Approx. Inoculum(CFU) | Recovery | Growth |
| --- | --- | --- | --- |
| Escherichia coli ATCC25922 | 10-100 | PR≥0.5 | Purple-red colonies |
| Enterococcus faecalis ATCC29212 |  |  | Turquoise colonies |
| Citrobacter fischeri ATCC43864 |  |  | Metallic blue with red halo colonies |
| Staphylococcus saprophyticus FSCC(I)223042 |  |  | Light pink colonies |
| Staphylococcus aureus ATCC 6538 |  |  | White to gold colonies |
| Streptococcus agalactiae ATCC12386 |  |  | Sapphirine colonies |
| Pseudomonas aeruginosa ATCC27853 |  |  | Translucent colonies,may be a faint green pigment |
| Proteus mirabilis CMCC(B)49005 |  |  | Brown halo colonies |

Storage and Shelf Life

Dry powder was stored in a sealed container at 2-30℃ for two years in a dry environment. The prepared plates were stored at 2-8℃ for one week ,Use before expiry date on the label.

Precautions

1. When weighing the dehydrated medium, please wear masks to avoid causing respiratory system discomfort

2. Keep container tightly closed after using to prevent clumping.

Waste Disposal

Microbiological contamination was disposed by autoclaving at 121°C for 30 minutes.

Revision

December1, 2024

Escherichia coli: Purple-red colonies             Enterococcus faecalis: Blue-green colonies

Citrobacter fischeri：Metallic blue with red halo colonies    Staphylococcus saprophyticus：Light pink colonies

Staphylococcus aureus：White to gold colonies             Streptococcus agalactiae：Sapphirine colonies

Pseudomonas aeruginosa：                    Proteus mirabilis：Brown halo colonies

Translucent colonies,may be a faint green pigment